KANCHI KAMAKOTI CHILDS TRUST HOSPITAL (KKCTH),

(A unit of “THE CHILDS TRUST”)

NO. 12A, NAGESWARA ROAD, NUNGAMBAKKAM, CHENNAI – 600 034.

Phone: 044-4200 1800, EXTN : 183

E-mail : matmanager@kkctch.org

GSTIN : 33AAAAT2552PIZA

Ref.No : KKCTH/ENQ/15/2025	                                      		Dt.  24.12.2026

INVITATION FOR QUOTATIONS FOR SUPPLY OF MEDICAL EQUIPMENT

THIS INVITATION DULY SIGNED ON ALL PAGES SHOULD BE ATTACHED WITH THE QUOTE

From

The Manager – Material Management

Kanchi Kamakoti CHILDS Trust Hospital (KKCTH)

Chennai – 600 034.

To

_____________________________

_____________________________
	_____________________________
Dear Sirs,

Sub:	Invitation for Quotations for supply and Installation of Medical Equipment to KKCTH, Chennai – Reg.

- You are invited to submit your most competitive quotation for the following equipment for the scope of supply and services indicated:-

| Sl. No. | Brief Description of the Equipment | Qty. in Nos. | Delivery Period | Place of Delivery |
| --- | --- | --- | --- | --- |
| 1. | Patient Monitor | 1 | Within 15 days from the date of receipt of supply order | Indicated at Annexure-II |

2.	Bid Price

2.1	The scope of work should include supply, transit insurance, unloading and storage at site, installation and commissioning at site.

2.2   The contract shall be for the full quantity as described above.  Corrections, if any, shall be made by crossing out, initialing, dating and re writing.

2.3	All duties, taxes payable by the contractor under the contract shall be considered in the quoted price indicating the breakup value.

2.4	a) If the items are imported, the bidder should furnish the CIF value of the imported component with Customs duty separately in the price bid in addition to the other breakup value  packaging & Forwarding, Inland transport, Installation and commissioning and other incidental charges

b) The bidders are advised to take in to consideration the 
savings available due to abolition of CVD and SAD for 
imports now due to GST implementation, and consider the 
benefits out of it as a reduction, while arriving at their bid 
price.

c) Similarly, the benefits that could be availed out of High Sea Sales/ Sales in course of import should also be considered as a reduction while arriving at their bid price.

d) GST shall be indicated separately.

2.5	The Prices should be quoted in Indian Rupees only.

3.	Each bidder shall submit only one quotation.

3.1. Earnest money deposit:-

  - Earnest Money Deposit :- Each quoted should be submitted with an EMD of Rs.1,000/- by means of DD drawn in favour of “Kanchi Kamakoti CHILDS Trust Hospital”, Chennai payable at Chennai.

  - Earnest Money Deposit in any other form such as Bank Guarantee / Cheque / Cash / Postal order will not be accepted.

3.2 Quotation received without EMD amount will be treated as non-    
      responsive.

3.3 The bidder should furnish the GST registration no. for supply and services and the code no. for the goods quoted.

3.4 The GST registered bidders are only eligible to participate in the quotation.

3.5 The bidder should furnish the details of Bank name, Branch name, Account no., IFSC code and a copy of cancelled cheque leaf.

4.	Validity of Quotation:-

Quotation shall remain valid for a period not less than 90 days after the deadline date specified for submission.

5.     Evaluation of Quotations:-

5.1	The Purchaser will evaluate and compare the quotations determined to be substantially responsive i.e. which

(a) are properly signed; and

- conform to the terms and conditions, and  specifications

- Not furnishing Catalogues, Point wise compliance statement of Technical specification, and producing the sample of the quoted item for demo within one week notice, if called for would be treated as a violation of the terms of the tender and the bids would be treated as non-responsive.

- The Quotation would be evaluated separately for each item considering the price of the equipment with accessories and Net Present Value (NPV) of the comprehensive AMC price at a discount rate of 8% per year as illustrated below:

| Sl. No. | Description | Unit rate (in Rs.) |
| --- | --- | --- |
| 1 | Equipment price with GST | X |
| 2 | CAMC price for the 2nd year after 1 year warranty period | 0.9259 X CAMC quoted Value (1/1.08=0.9259) |
| 3 | CAMC price for the 3nd year | 0.8573 X CAMC quoted Value (1/(1.08^2)) = 0.8573 |
| 4 | CAMC price for the 4th year | 0.7938 X CAMC quoted Value (1/(1.08^3)) = 0.7938 |
| 5 | CAMC price for the 5nd year | 0.7350 X CAMC quoted Value (1/(1.08^4)) = 0.7350 |
| 6 | CAMC price for the 6th year | 0.6805 X CAMC quoted Value 1/(1.08^5)) = 0.6805 |
| 7 | Total CAMC Discounted Value (Incl. GST) | (2+3+4+5+6) |
| 8 | Evaluated Price with GST + NPV of CAMC with GST | (X+7) |

6.	Award of contract:-

The Purchaser will award the contract to the bidder whose quotation has been determined to be substantially responsive and who has offered the lowest evaluated quotation price.

6.1	Notwithstanding the above, the Purchaser reserves the right to accept or reject any quotations and to cancel the bidding process and reject all quotations at any time prior to the award of contract.

6.2	The bidder whose bid is accepted will be notified of the award of contract by the Purchaser prior to expiration of the quotation validity period.  The terms of the accepted offer shall be incorporated in the purchase order.

6.3	For delay in supply of the goods beyond the stipulated delivery period, liquidated damages at 0.5% per week or part thereof of undelivered portion of the contract, subject to a maximum of 10% of the contract value is leviable.

6.4	The successful bidder shall furnish performance security for 5%

of contract value by means of demand draft drawn in favour of

Kanchi  Kamakoti CHILDS Trust Hospital, Chennai Payable at

Chennai which will be returned after satisfactory completion of

warranty period.

6.5	Payment shall be made within 30 days after delivery of the goods, against submission of bills with certification from the consignee for satisfactory completion of supply.

6.6	The item supplied shall carry a warranty of atleast 1 year for 
	equipment from the date of acceptance of the goods by the 
	consignee.

6.7 Annual Maintenance Charges (Comprehensive) should be 
quoted for 5 years / per year after 1 year warranty maintenance 
 period for equipment which is extendable for further period on mutual consent.

7.  You are requested to submit the quotations in a sealed cover superscribed “Quotations for Patient Monitor to Kanchi Kamakoti CHILDS Trust Hospital, Chennai” and quotations should reach Kanchi Kamakoti CHILDS Trust Hospital (No. 12A, Nageswara Road, Nungambakkam, Chennai – 600 034) on or before 11.00 AM on 9th Jan 2026.

8. Quotation will be opened at 11.30 AM on 9th Jan 2026 in the presence of bidder rep. who chose to be present and at the sole discretion of purchaser.

9. We look forward to receive your quotations and thank you for your interest in this project.

Manager – Material Management

If any Query, please contact our Material Manager

Mail id : matmanager@kkcth.org , bmm@kkcth.org

Cc:  childsstrustfrandim@kkcth.org

PRICE SCHEDULE

| 1 | 2 | 3 | 4 | 5 |  |  |  |  | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | PRICE FOR EACH UNIT |  |  |  |  |  |  |  |
| Sch No. | Item Description | Country of origin | Quantity & Unit | Ex-factory Ex-warehouse Ex-showroom off-the shelf for CIF (a) | Packing & forwarding (b) | Inland transport, Insurance and Incidental costs incidental to delivery € | Incidental services (d) | Customs duty (e) | Unit price a+b+c+d+e | Total price per schedule for delivery at final destination (4 x 6) | GST – IGST/CGST/ SGST payable if contract is awarded |
| 1. | Patient Monitor as per specification |  | 1 |  |  |  |  |  |  |  |  |

i. Unit price in (6) (Rs. in words)

1. Patient Monitor as per specification with HSN code/ rate of GST……………………………………….……………………………………………………………….

a) Annual Maintenance Charges (Comprehensive) for 5 years / per year / per unit after free warranty maintenance period

1. Patient Monitor as per specification with HSN code/ rate of GST.................................................................................................................................................................

Note:

(a)	In case of discrepancy between unit price and total price, the unit price shall prevail.

(b)	GST applicable for Annual Maintenance Charges shall be indicated separately.

(c)  	The bidder should indicate the HSN code of the equipment/ service and applicable GST rate

(d)	The bidder should quote for the prices separately from column 5(a) to 5(e) and should not state “as inclusive” of all rate.

Place:									Signature of Bidder......................……………..

Date:									Name                      ……………………………

Business Address……………………………

ANNEXURE I

TECHNICAL SPECIFICATION FOR PATIENT MONITOR

GENERAL

Product Description					High End Multipara Monitor

Purpose	High end multipara monitor is a device that monitor and record a variety of patients' basic vital signs parameters in real-time. These are used in the critical care unit, Operation theatres and in post anesthesia care unit and ICU.

PRODUCT INFORMATION

Type of Monitor						Modular

Monitoring Parameters	ECG, HR, Respiration Rate, SpO2, Pulse Rate, NIBP, Temperature, IBP

EtCO2	Upgradable

Cardiac output	NA

Screen size of transport module in inches	NA

Battery backup for transport monitor in hrs. 	NA

AGM (Anesthetic Gas Module)	NA

Channels in EEG	NA

Perfusion Index 	Yes

PPV (derived value during IBP measurement)	Yes

Length of time of full disclosure 	≥ 24

of patient data in hr.

Length of time of trending in hr. 	≥ 24

ECG PARAMETERS

ECG Provided		Yes

Number of leads of ECG		5

12-Lead interpretation for ECG		Yes

ECG Waveform number of leads display		12

Arrhythmia detect		Yes

Heart rate measurement range		Atleast 30 -250 bpm

Accuracy		± 2 bpm or better

Minimum gradation		1 bpm or better

RESPIRATION PARAMETERS

Respiration Method						Impedance pneumography

Display waveform of respiration				Yes

Respiration rate measurement range				0 - 100 bpm or more

Minimum gradation of respiratory measurement		1 bpm or better

PULSE OXIMETER PARAMETERS

Measurement technique (Declaration of the			Low Perfusion Technology with

measurement technique by the seller at the			motion artifact detection

time of bid submission/supply)

Probe type								Rubber

Probe for Adult and Pediatric					Reusable

Probe for neonates						Disposable

SPO2 measurement range					70 - 100%

Accuracy of SPO2 measurement				± 2% or better

Minimum Gradation						1%

NIBP PARAMETERS

Measurement technique for NIBP	Step-deflation Oscillometric / Linear Inflation

Systolic Blood pressure monitoring range	40 - 250 mmHg or more

Diastolic Blood pressure monitoring range	30 - 220 mmHg or more

Accuracy of NIBP Measurement	±3 mm/Hg or better

TEMPERATURE PARAMETER

Number of channels for temperature	2

Type of probe	Skin, Rectal, Esophageal

Temperature probe type	Reusable

Temperature probe range	Atleast 30 - 40 degree

Minimum gradation of temperature probe	0.1 degree C or better

Accuracy in %	±1% or better

IBP PARAMETERS

Number of channels for IBP	2

ETCO2 PARAMETER

Patient type	Non-intubated

ALARMS

Availability of Audio and visual alarms 		Yes

with high, medium and low intensity

adjustment.

Operator can set audio visual alarm levels 		Yes

for low or high levels of each parameter

Independently

Systolic BP - low and high		Yes

Diastolic BP - low and high		Yes

Mean BP - low and high		Yes

Heart rate - low and high		Yes

SPO2 - low and high		Yes

Respiratory rate - low and high		Yes

Temperature - low and high		Yes

Alarm Volume Adjustable		Yes

Activating silence of alarm		Yes

Duration of alarm silence in seconds		300

Suspension of all alarms		Yes

Latching of alarm		Yes

Cuff leakage alarm		Yes

Cuff disconnect alarm		Yes

Alarm for inflation/deflation errors		Yes

Indicator for low-battery		Yes

EtCO2 - low and high		Yes

Escalation of Alarm Priority		Yes

Sensor/ wire/probe disconnection alarm		Yes

NETWORK PARAMETERS

Ethernet port	Yes

Hardwired	Yes

Connectivity with Network printer	Yes

Connectivity provision with Central Nursing Station	Yes

DISPLAY PARAMETERS

Type of display	LCD/LED/TFT

Display Size in inches	12

Traces of display	>6

User Interface of display	Touchscreen

STANDARD ACCESSORIES

Number of reusable adult probe provided	1

Number of reusable pediatric probe provided	1

Number of ECG 5/6/10 lead cable provided	1

Number of Neonatal ECG 3/5/6 lead cable provided	1

Number of IBP cables provided	2 Reusable Cables

Number of Adult NIBP cuffs of Medium,	1 Set

Large, Extra large provided

Number of Pediatric NIBP cuffs of Small 	1 Set

and medium provided

Number of Skin temperature probe provided	1 for Adult / Paediatric

Accessories to wall mount the monitor	Yes

POWER REQUIREMENTS

Power supply	220-240 VAC, 50Hz , Single phase

Mains cable	3 meter or more with Indian plug

Battery	Inbuilt

Minimum Inbuilt Battery back up in minutes	60

Recharging time of battery in hr	≤ 6

Weight of the Machine in kgs	≤ 15

Noise level in dB	≤ 50

Device Earthing	Yes

CERTIFICATIONS

Compliance to Medical Device Rules	Yes

(MDR) 2017 as amended till date

Availability of valid Medical Device license for the 	Yes

product issued from the competent authority

defined under Drugs and Cosmetic Act 1940

and Rules made there under as amended

till date

Valid Medical Device License Number	HQ/MD/2022/002360

Certification for manufacturing unit	ISO:13485 (Latest)

Availability of Test Report for each supplied 	Yes

batch/product as per Medical Device Rule

(MDR) 2017 as amended till date

Submission of all necessary certifications	Yes

licenses and test reports to the buyer at the

time of bid submission or along with

supplies as per buyer requirement

Conformity to Electrical Safety Standards 	IEC 60601 or BIS Equivalent

Note:

  - Bidders shall furnish technical compliance statement for the model quoted, details of manufacturer including deviations if any. Technical catalogue /data sheet shall also be furnished in support of technical compliance statement without fail.

  - Not demonstrating the equipment when requested by purchaser is a reason for rejection.

  - “Any breakdown resulting from a rodent issue will not be borne by the end user” and that it will be the responsibility of the supplier to prevent the entry of rodent inside the equipment.

Annexure II

Place of Delivery

| Sl. No. | Name of the hospital | Qty (in nos.) |
| --- | --- | --- |
| 1 | Kanchi Kamakoti CHILDS Trust Hospital, Chennai | 1 |