KANCHI KAMAKOTI CHILDS TRUST HOSPITAL (KKCTH),

(A unit of “THE CHILDS TRUST”)

NO. 12A, NAGESWARA ROAD, NUNGAMBAKKAM, CHENNAI – 600 034.

Phone: 044-4200 1800, EXTN : 183

E-mail : matmanager@kkctch.org

GSTIN : 33AAAAT2552PIZA

TENDER FOR SUPPLY OF MEDICAL EQUIPMENT

TENDER NO : KKCTH/ENQ/16/2025, Dt. 24.12.2026

THIS DOCUMENT DULY SIGNED ON ALL PAGES SHOULD BE ATTACHED WITH THE BID

From

The Manager – Material Management

Kanchi Kamakoti CHILDS Trust Hospital (KKCTH)

Chennai – 600 034.

To

_____________________________

_____________________________
	_____________________________
Dear Sirs,

Sub:	Tender for supply and Installation of Medical Equipment to KKCTH, Chennai – Reg.

- INVITATION OF BID

You are invited to submit your tender with most competitive rates for the following equipment conforming to the technical specification and other terms and conditions and scope of supply and services indicated:-

| Sl. No. | Brief Description of the Equipment | Qty. in Nos. | Delivery Period | Place of Delivery |
| --- | --- | --- | --- | --- |
| 1. | Pediatric ICU Ventilator as per specification | 2 | Within 15 days from the date of receipt of supply order | Kanchi Kamakoti CHILDS Trust Hospital, Chennai |

2. ELIGIBILITY CRITERIA

2.1 The bidder should either be manufacturer or non-manufacturer with valid manufacturer authorization to quote this product against this tender.	 Copy of the manufacturer authorization indicating the tender no. and date mentioned above should be furnished.

2.2 The bidder as manufacturer or non-manufacturer bidder should have experience of manufacture and supply and providing after sales service of similar equipment and of same make for atleast in the last three years.

2.3 The manufacturer should have supplied either directly or through any other dealer atleast 1 no. of similar model of their make in any one of the last five years with satisfactory performance of atleast one year on the date of bid opening supported by Purchase order, Installation report and End user certificate.

2.4 The non-manufacturer bidder in addition to their manufacturer meeting the above condition should have supplied and provided after sales service atleast 1 no. of similar model of the same make in any one of the last five years with satisfactory performance of atleast one year on the date of bid opening supported by Purchase order, Installation report and End user certificate.

2.5 The bidder should have an annual turnover of Rs. 50.00 lakhs in any one of the last three years, supported by a Chartered Accountant certificate, and the bidder should have a cash credit facility of at least Rs. 10.00 lakhs, certified by their bankers.

2.6 The bidder should not have blacklisted for any of their previous supply to private or public organization and blacklisting should not be in force on the date of bid opening. An undertaking in this regard should be furnished by the bidder.

3. BID PRICE

3.1	The scope of work should include supply, transit insurance, unloading and storage at site, installation and commissioning at site.

3.2   The contract shall be for the full quantity as described above.  Corrections, if any, shall be made by crossing out, initialing, dating and re writing.

3.3	All duties, taxes payable by the contractor under the contract shall be considered in the quoted price indicating the breakup value.

3.4	a) If the items are imported, the bidder should furnish the CIF value of the imported component with Customs duty separatelyn the price bid in addition to the other breakup value  packaging & Forwarding, Inland transport, Installation and commissioning and other incidental charges

b) The bidders are advised to take in to consideration the 
savings available due to abolition of CVD and SAD for 
imports now due to GST implementation, and consider the 
benefits out of it as a reduction, while arriving at their bid 
price.

c) Similarly, the benefits that could be availed out of High Sea Sales/ Sales in course of import should also be considered as a reduction while arriving at their bid price.

d) GST shall be indicated separately.

3.5	The Prices should be quoted in Indian Rupees only.

4.	NO. OF BIDS AND EARNEST MONEY DEPOSIT:

4.1 Each bidder shall submit only one bid.

  - Earnest Money Deposit :- Each quoted should be submitted with an EMD of Rs.1,000/- by means of DD drawn in favour of “Kanchi Kamakoti CHILDS Trust Hospital”, Chennai payable at Chennai.

  - Earnest Money Deposit in any other form such as Bank Guarantee / Cheque / Cash / Postal order will not be accepted.

4.2 Tender received without EMD amount will be treated as non-    
      responsive.

4.3 The bidder should furnish the GST registration no. for supply and services and the code no. for the goods quoted.

4.4 The GST registered bidders are only eligible to participate in the tender.

4.5 The bidder should furnish the details of Bank name, Branch name, Account no., IFSC code and a copy of cancelled cheque leaf.

5.	VALIDITY OF THE BID:-

The bid shall remain valid for a period not less than 90 days after the deadline date specified for submission.

6.     EVALUATION OF THE BIDS:-

6.1	The Purchaser will evaluate and compare the price bids of the tender determined to be substantially responsive i.e. which

(a) are properly signed; and

- conform to the terms and conditions, and specifications

- Not furnishing Catalogues, Point wise compliance statement of Technical specification, and producing the sample of the quoted item for demo within one week notice, if called for would be treated as a violation of the terms of the tender and the bids would be treated as non-responsive.

- The prices will be evaluated separately for each item considering the price of the equipment with accessories and Net Present Value (NPV) of the comprehensive AMC price at a discount rate of 8% per year as illustrated below:

| Sl. No. | Description | Unit rate (in Rs.) |
| --- | --- | --- |
| 1 | Equipment price with GST | X |
| 2 | CAMC price for the 2nd year after 1 year warranty period | 0.9259 X CAMC quoted Value (1/1.08=0.9259) |
| 3 | CAMC price for the 3nd year | 0.8573 X CAMC quoted Value (1/(1.08^2)) = 0.8573 |
| 4 | CAMC price for the 4th year | 0.7938 X CAMC quoted Value (1/(1.08^3)) = 0.7938 |
| 5 | CAMC price for the 5nd year | 0.7350 X CAMC quoted Value (1/(1.08^4)) = 0.7350 |
| 6 | CAMC price for the 6th year | 0.6805 X CAMC quoted Value 1/(1.08^5)) = 0.6805 |
| 7 | Total CAMC Discounted Value (Incl. GST) | (2+3+4+5+6) |
| 8 | Evaluated Price with GST + NPV of CAMC with GST | Y= (X+7) |
| 9 | Cost of cartridge for 100 nos. with GST | Z |
| 10 | Grand evaluated price for L1 assessment | Y+Z |

7.	AWARD OF CONTRACT:-

The Purchaser will award the contract to the bidder whose tender has been determined to be substantially responsive and who has offered the lowest evaluated price.

7.1	Notwithstanding the above, the Purchaser reserves the right to accept or reject any bids and to cancel the bidding process and reject all the bids at any time prior to the award of contract.

7.2	The bidder whose bid is accepted will be notified of the award of contract by the Purchaser prior to expiration of the tender validity period. The terms of the accepted offer shall be incorporated in the purchase order.

7.3	For delay in supply of the goods beyond the stipulated delivery period, liquidated damages at 0.5% per week or part thereof of undelivered portion of the contract, subject to a maximum of 10% of the contract value is leviable.

7.4	The successful bidder shall furnish performance security for 5%

of contract value by means of demand draft drawn in favour of

Kanchi  Kamakoti CHILDS Trust Hospital, Chennai Payable at

Chennai which will be returned after satisfactory completion of

warranty period and submission of another bank guarantee for

5% value of total CAMC value valid till the end of CAMC period.

7.5	Payment shall be made within 30 days after delivery of the goods, against submission of bills with certification from the consignee for satisfactory completion of supply.

7.6	The item supplied shall carry a warranty of atleast 1 year for 
	equipment from the date of acceptance of the goods by the 
	consignee.

8. ANNUAL MAINTENANCE

8.1 Annual Maintenance Charges (Comprehensive) should be 
quoted for 5 years / per year after 1 year warranty maintenance 
 period for equipment which is extendable for further period on mutual consent. The successful bidder should make one visit per quarter under preventive maintenance and any no. of visits for breakdown calls. Uptime guarantee of 95% should be assured in a year. The breakdown calls should be attended and rectified within 48 hours of call. If for reasons beyond the control rectification could not be done within 48 hours, then a standby machine should be placed at free of cost till the original machine is rectified and handed over. Any failure in the above will attract a penalty of Rs. 5,000/- per day.

9. DISPUTE RESOLUTION

9.1 Any contractual dispute shall be resolved on mutual discussion and if it could not be resolved, it shall be referred to arbitration under Arbitration and Conciliation Act. The jurisdiction is Tamil Nadu

10.  BID SUBMISSION

10.1 The bid should be submitted in two cover system namely Technical bid and Price bid in separate sealed covers which are placed in an outer cover and each cover superscribed “Tender for Pediatric ICU Ventilator to Kanchi Kamakoti CHILDS Trust Hospital, Chennai” and tender should reach Kanchi Kamakoti CHILDS Trust Hospital (No. 12A, Nageswara Road, Nungambakkam, Chennai – 600 034) on or before 11.00 AM on 9th Jan 2026.

11. OPENING OF BIDS

11.1 The technical bids will be opened at 11.30 AM on 9th Jan 2026 in the presence of bidder rep. who chose to be present and at the sole discretion of purchaser. After evaluation of the technical bids, the price bids of the responsive bidders will be opened at a later date under prior intimation to the responsive bidders. The price bids of the non-responsive bidders will not be opened and they are not permitted to attend the price bid opening.

12. We look forward to receive your bid and thank you for your interest in this project.

Manager – Material Management

If any Query, please contact our Material Manager

Mail id : matmanager@kkcth.org , bmm@kkcth.org

Cc:  childstrustfrandim@kkcth.org

PRICE SCHEDULE

| 1 | 2 | 3 | 4 | 5 |  |  |  |  | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | PRICE FOR EACH UNIT |  |  |  |  |  |  |  |
| Sch No. | Item Description | Country of origin | Quantity & Unit | Ex-factory Ex-warehouse Ex-showroom off-the shelf for CIF (a) | Packing & forwarding (b) | Inland transport, Insurance and Incidental costs incidental todelivery € | Incidental services (d) | Customs duty (e) | Unit price a+b+c+d+e | Total price per schedule for delivery at final destination (4 x 6) | GST – IGST/CGST/ SGST payable if contract is awarded |
| 1. | Pediatric ICU Ventilator as per specification |  | 2 nos. |  |  |  |  |  |  |  |  |

i. Unit price in (6) (Rs. in words)

1. Pediatric ICU Ventilator as per specification with HSN code/ rate of GST……………………………………….

a) Annual Maintenance Charges (Comprehensive) for 5 years / per year / per unit after free warranty maintenance period

1. Pediatric ICU Ventilator as per specification with HSN code/ rate of GST.....................................................................

Note:

(a)	In case of discrepancy between unit price and total price, the unit price shall prevail.

(b)	GST applicable for Annual Maintenance Charges shall be indicated separately.

(c)  	The bidder should indicate the HSN code of the equipment/ service and applicable GST rate

(d)	The bidder should quote for the prices separately from column 5(a) to 5(e) and should not state “as inclusive” of all rate.

Place:									Signature of Bidder......................……………..

Date:									Name                      ……………………………

Business Address……………………………

ANNEXURE I

TECHNICAL SPECIFICATION FOR PEDIATRIC ICU VENTILATOR

GENERAL

| Product Description | Pediatric ICU Ventilator |
| --- | --- |
| Purpose | Intensive care ventilators are defined as mechanical ventilators that can be configured to provide invasive ventilation (e.g., with an endotracheal tube or tracheostomy tube) and noninvasive ventilation (eg, with a face mask) |

PRODUCT INFORMATION

| Patient Type | Neonatal, Pediatric and Adult |
| --- | --- |
| Tidal volume in ml | 2 ml (or less) to 4,000 ml (or more) |
| Respiration rate, breaths/min | upto 150 or more |
| Trigger mechanism | Both pressure and flow; upgradable to neural trigger (NAVA) |
| FiO2% | 21 to 100 |
| Inspiratory flow rate, L/min | upto 198 L/min or more |
| Inspiratory pressure, cm H2O | upto 100 cm H2O or more |
| IE ratio | 1:10 to 4:1 |
| Sigh breath function | Yes |
| PEEP/CPAP, cm H2O | upto 40 or more |
| Pressure support, cm H2O | 0 to 40 or more |
| Leak compensation | Yes |
| Auto 100%/Increase O2 button | Yes |
| Control panel lock | Yes |
| Flow and volume measurements | Shall be made by advanced ultrasonic technology unsusceptible to moisture and patient secretions. |
| Inspired oxygen concentration measurement | Shall be made Lead free, non-consumable technology with long life time of minimum 10 years |

PATIENT ASSESSMENT TOOLS

| Minimum wave forms displayed | 5 (with the related numerical parameters) |
| --- | --- |
| Number of loops | 2 loops (P-V, F-V) with facility of saving of 1 Loops for reference |
| Lung recruitment tools (PV loops) / OLT (Open lung tools) | Optional - Upgradable |
| Lung mechanics visualization tool | Yes |
| Capnography/CO2 monitoring | Optional - Upgradable |
| Esophageal / transpulmonary pressure monitoring | Optional - Upgradable |
| Modes of ventilation | Volume control, Pressure control, Pressure Regulated Volume Control, SIMV (VC, PC, PRVC) +PS, Pressure support, Volume Support, Bi-Vent, CPAP / PEEP, NIV, Nasal CPAP and High Flow Therapy |
| Advanced Modes: NAVA (Neurally Adjusted Ventilatory Assist) | Optional - Upgradable |
| Ventilation monitoring facility | The Ventilator is a Microprocessor Controlled ventilator with integrated facility for Ventilation monitoring |
| User interface | The user interface shall have media library, which enables export of screenshots and recordings to a USB-stick. |
| Monitored / Displayed parameters | Peak inspiratory pressure, Mean airway pressure, PEEP pressure, Tidal volume, Minute volume |
| Availability of Status indicator on display | Ventilator mode, Battery life, patient data, alarm settings, clock |
| Patient alarms | Low/high FiO2,Low/high minute volume, Low inspiratory pressure, High pressure, Low PEEP, High PEEP, Apnea |
| Equipment alarms | Gas-supply failure, Power failure, Vent inoperative, Low battery, Self-diagnostic |
| Expiratory Time constant ( Tcexp) | Yes |
| Compliance static and/or dynamic | Yes |
| Inspired and/or expired resistance | Yes |
| Occlusion pressure | Yes |
| Inspiratory and expiratory hold | Yes |
| Spontaneous frequency | Yes |
| Total peep, intrinsic peep, extrinsic peep | Yes |
| Driving pressure display | Yes |
| Plateau/Driving Pressure targets | It shall be possible to see when plateau/ driving pressure or tidal volume per predicted body weight (VT/PBW) off pre-defined targets and interventions are needed. |
| Automatic Stepwise Recruitment Maneuver | Optional - Upgradable |
| Volumetric capnography | Optional - Upgradable |
| Expiratory block is Autoclavable and no routine calibration required | Yes |
| Facility for automatic compliance & Leakage compensation for circuit available | Yes |
| Patient safety systems | The system shall include a mode, which allows patients to go into a support mode automatically if they trigger the ventilator and If the patient is not making any breathing effort the ventilator will deliver controlled breaths. |
| Facility to set Inspiratory tidal volume in relation to predicted body weight | Yes |
| RS 232 out put port | Yes |
| VGA-port for connection to slave screen/projector | Yes |
| USB/Memory card | Yes |
| Facility to Report (vent alarms and patient status) | Yes |
| HL/7 Compliant out put | Yes |
| Type of coloured display | LED/LCD/TFT |
| Display should be touch screen | Yes (capacitive touch screen which also supports use with gloves) |
| Minimum Size of display (±2%) | 15 inch, with 360 degrees rotation facility |
| Graphic display has automatic scaling facility for waves | Yes |
| Remote service function | The system shall be equipped with a remote service function with direct access connection, without use of an additional docking station. |

HARDWARE AND ELECTRICAL REQUIREMENTS

| Power supply | 220-240 V , 50 Hz AC |
| --- | --- |
| Internal backup battery | Yes – hot swappable NiMH type |
| Backup time of internal battery in mins | 2 hours or more |
| Integrated control for micropump Nebulizer for intermittent and continuous nebulization with particle size less than 5 micron | Yes |

ADDITIONAL ACCESSORIES

| Number of reusable silicon breathing circuits | 2 |
| --- | --- |
| Length of Power cord in m | ≥ 3 m |
| Autoclavable exhalation valve/expiratory cassette / filter | 2 |
| Trolley for ventilator with circuit holding arm | Yes |

CERTIFICATIONS

| Compliance to Medical Device Rules (MDR) 2017 as amended till date | Yes |
| --- | --- |
| Availability of valid Medical Device license for the product issued from the competent authority defined under Drugs and Cosmetic Act 1940 and Rules made there under as amended till date | Yes |
| Valid Medical Device License Number | To be Furnished |
| Certification for manufacturing unit | ISO:13485 (Latest) |
| Availability of Test Report for each supplied batch/product as per Medical Device Rule (MDR) 2017 as amended till date | Yes |
| Submission of all necessary certifications, licenses and test reports to the buyer at the time of bid submission or along with supplies as per buyer requirement | Yes |
| Alarm priority | Alarm priority in terms of sound and color-coding shall meet IEC 60601-1-8:2006; It shall be possible to visually detect alarms from an angle of 360° around the system |
| Standards to be complied with | The ventilator system should fulfill the 3rd edition medical devices standards for mechanical ventilators, i.e. IEC 60601-1:2005, ISO 80601-2-12:2011, ISO 80601-2-55:2011 and shall have Ingress protection (IP21). |

Note:

  - Bidders shall furnish technical compliance statement for the model quoted, details of manufacturer including deviations if any. Technical catalogue /data sheet shall also be furnished in support of technical compliance statement without fail.

  - Not demonstrating the equipment when requested by purchaser is a reason for rejection.

Annexure II

Place of Delivery

| Sl. No. | Name of the hospital | Qty (in nos.) |
| --- | --- | --- |
| 1 | Kanchi Kamakoti CHILDS Trust Hospital, Chennai | 2 |